## Summary
pH validation — summary of runs
All measurements at fixed 1:2.5 ratio in 0.01 M CaCl₂. Reference soil established value: 8.03 ± 0.11 (5 g).
Instruments used
Probe | Lab | Type | Slope | Condition | Verdict
Probe 1 | Collaborating lab (Prof. Magnus) | Microelectrode | 92% | Cracked glass, low slope | Faulty — readings not trusted
Probe 2 | Collaborating lab (Prof. Magnus) | Semi-micro | 97% | Healthy slope | Reads ~0.28 low vs reference
Probe 3 | Our lab | Regular combination (accumet AB315) | 92.9–98.8%* | pH 10 check: 10.07 / 10.06 / 10.08 | Validated — used for final data
* The meter's slope field displayed inconsistently without recalibration; the pH 10 check standard was used as the operative validity criterion.
Timeline
Date | Lab / probe | What was done
7 July 2026 | Magnus lab, Probe 1 (micro) | Reference soil mass series (1 / 2 / 5 g), triplicate. Probe later found cracked — absolute values not trusted.
8 July 2026 | Magnus lab, Probe 2 (semi-micro) | Reference soil + A1, A8 at 1 / 2 / 5 g. Readings low; leftover vs fresh CaCl₂ tested to rule out the electrolyte.
9 July 2026 | Our lab, Probe 3 (regular) | Reference soil validated at 8.03 (50 mL, triplicate); 15 mL tube check; 1 g vs 5 g mass comparison.
Key findings
• Reference soil on our validated probe (5 g): mean 8.03, SD 0.05 — matches established 8.03 ± 0.11.
• The semi-micro probe (Magnus lab) reads ~0.28 low on the same reference soil — an instrument offset, not a sample effect.
• CaCl₂ ruled out: fresh 0.01 M did not restore the expected value on the low-reading probe.
• Vessel geometry ruled out: 5 g gave 8.03 (50 mL tube) and 8.08 (15 mL tube).
• 1 g bias: with immersion resolved, 1 g reads ~0.15 below 5 g (7.93, SD 0.00). The same downward direction appeared on 7 July (−0.28), on a different electrode.
• Biological replicates (~0.15) scatter more than technical re-reads (~0.05): subsampling dominates, not the meter.
Proposed next step
Samples are limited to 1 g. Rather than carrying a correction factor through the dataset, characterise a new reference batch directly at 1 g on our validated electrode (~10 fresh subsamples across ≥3 days and ≥2 calibrations), running a few 5 g scoops alongside to confirm the mass offset independently. The previously characterised batch is exhausted.
Empty Quarter — pH measurement log (by trip)
• Sample counts — Trip 1: 330, Trip 2: 27, Trip 3: 450, Trip 4: 177, Trip 5: 234.

## Trip 1
Sample ID | Site | Compartment | Rep | Date | Calibration | Slope % | pH 10 check | pH reading | Notes
1Sr1 | 1 | Surface | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.99
1Sr2 | 1 | Surface | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 8
1Sr3 | 1 | Surface | 3 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.98
1Dr1 | 1 | Deep | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 8
1Dr2 | 1 | Deep | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.99
1Dr3 | 1 | Deep | 3 | 2026-07-21 | Yes | 97.4 | 10.06 | 8
1PRr2 | 1 | Rhizosphere | 2
1PRr3 | 1 | Rhizosphere | 3 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.95
2Sr1 | 2 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.15
2Sr3 | 2 | Surface | 3
2Dr1 | 2 | Deep | 1
2Dr2 | 2 | Deep | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.07
2Dr3 | 2 | Deep | 3
2PRr1 | 2 | Rhizosphere | 1
2PRr2 | 2 | Rhizosphere | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.9
3Sr1 | 3 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.84
3Sr3 | 3 | Surface | 3
3Dr2 | 3 | Deep | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.77
3Dr3 | 3 | Deep | 3
3PRr1 | 3 | Rhizosphere | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.65
3PRr2 | 3 | Rhizosphere | 2
3PRr3 | 3 | Rhizosphere | 3
4Sr1 | 4 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.02
4Dr1 | 4 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.97
4Dr2 | 4 | Deep | 2
4Dr3 | 4 | Deep | 3
4PRr1 | 4 | Rhizosphere | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.93
4PRr2 | 4 | Rhizosphere | 2
4PRr3 | 4 | Rhizosphere | 3
5Sr1 | 5 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.03
5Sr2 | 5 | Surface | 2
5Sr3 | 5 | Surface | 3
5Dr1 | 5 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.99
5Dr2 | 5 | Deep | 2
6Sr1 | 6 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.96
6Sr2 | 6 | Surface | 2
6Dr1 | 6 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.93
6Dr3 | 6 | Deep | 3
6PRr1 | 6 | Rhizosphere | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.88
6PRr2 | 6 | Rhizosphere | 2
7Sr1 | 7 | Surface | 1
7Sr2 | 7 | Surface | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.04
7Dr1 | 7 | Deep | 1
7Dr2 | 7 | Deep | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.95
7PRr2 | 7 | Rhizosphere | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.7
8Sr2 | 8 | Surface | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.4
8Sr3 | 8 | Surface | 3 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.39
8PRr1 | 8 | Rhizosphere | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.71
8PRr2 | 8 | Rhizosphere | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.77
8PRr3 | 8 | Rhizosphere | 3 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.69
9Sr1 | 9 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.09
9Sr2 | 9 | Surface | 2
9Sr3 | 9 | Surface | 3
9Dr1 | 9 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.97
9Dr3 | 9 | Deep | 3
9PRr2 | 9 | Rhizosphere | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.9
9PRr3 | 9 | Rhizosphere | 3
10Sr2 | 10 | Surface | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.01
10Sr3 | 10 | Surface | 3
10Dr2 | 10 | Deep | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 7.7
10PRr2 | 10 | Rhizosphere | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.02
10PRr3 | 10 | Rhizosphere | 3
11Sr2 | 11 | Surface | 2 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.12
11Dr1 | 11 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.06 | 8.13
11Dr3 | 11 | Deep | 3
11PRr1 | 11 | Rhizosphere | 1
11PRr2 | 11 | Rhizosphere | 2
11PRr3 | 11 | Rhizosphere | 3 | 2026-07-22 | Yes | 96.1 | 10.07 | 8.01
12Sr1 | 12 | Surface | 1 | [5 empty] | Depleted
12Sr2 | 12 | Surface | 2 | [5 empty] | Depleted
12Sr3 | 12 | Surface | 3 | [5 empty] | Depleted
12Dr2 | 12 | Deep | 2 | [5 empty] | Depleted
12Dr3 | 12 | Deep | 3 | [5 empty] | Depleted
12PRr1 | 12 | Rhizosphere | 1 | [5 empty] | Depleted
12PRr3 | 12 | Rhizosphere | 3 | [5 empty] | Depleted
13Sr2 | 13 | Surface | 2 | 2026-07-22 | Yes | 96.1 | 10.07 | 8.02
13Dr1 | 13 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 8.03
13Dr3 | 13 | Deep | 3
14Sr2 | 14 | Surface | 2 | [5 empty] | Depleted
14Dr1 | 14 | Deep | 1 | [5 empty] | Depleted
14Dr2 | 14 | Deep | 2 | [5 empty] | Depleted
14Dr3 | 14 | Deep | 3 | [5 empty] | Depleted
14PRr1 | 14 | Rhizosphere | 1 | [5 empty] | Depleted
14PRr2 | 14 | Rhizosphere | 2 | [5 empty] | Depleted
15Sr1 | 15 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.97
15Sr3 | 15 | Surface | 3
15Dr1 | 15 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.93
15Dr3 | 15 | Deep | 3
15PRr1 | 15 | Rhizosphere | 1
15PRr2 | 15 | Rhizosphere | 2 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.8
15PRr3 | 15 | Rhizosphere | 3
16Sr1 | 16 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.88
16Dr1 | 16 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.91
16Dr3 | 16 | Deep | 3
16PRr1 | 16 | Rhizosphere | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.8
16PRr2 | 16 | Rhizosphere | 2
16PRr3 | 16 | Rhizosphere | 3
17Sr1 | 17 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 8.01
17Sr2 | 17 | Surface | 2
17Sr3 | 17 | Surface | 3
17Dr1 | 17 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.99
17Dr2 | 17 | Deep | 2
17Dr3 | 17 | Deep | 3
17PRr2 | 17 | Rhizosphere | 2 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.7
17PRr3 | 17 | Rhizosphere | 3
18Sr1 | 18 | Surface | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.82
18Sr2 | 18 | Surface | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.93
18Dr2 | 18 | Deep | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.61
18PRr1 | 18 | Rhizosphere | 1
18PRr2 | 18 | Rhizosphere | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.91
19Sr1 | 19 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.78
19Sr3 | 19 | Surface | 3
19Dr2 | 19 | Deep | 2 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.7
19Dr3 | 19 | Deep | 3
20Sr1 | 20 | Surface | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.81
20Sr3 | 20 | Surface | 3
20Dr1 | 20 | Deep | 1 | 2026-07-22 | Yes | 96.1 | 10.07 | 7.79
20Dr2 | 20 | Deep | 2
20PRr1 | 20 | Rhizosphere | 1
20PRr2 | 20 | Rhizosphere | 2
21Sr1 | 21 | Surface | 1
21Sr2 | 21 | Surface | 2
21Sr3 | 21 | Surface | 3
21Dr1 | 21 | Deep | 1
22Sr1 | 22 | Surface | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.98
22Sr3 | 22 | Surface | 3 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.91
22Dr1 | 22 | Deep | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.99
22Dr2 | 22 | Deep | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.96
22Dr3 | 22 | Deep | 3
22PRr1 | 22 | Rhizosphere | 1
22PRr3 | 22 | Rhizosphere | 3 | 2026-07-21 | Yes | 97.4 | 10.06 | 7.97
23Sr1 | 23 | Surface | 1
23Sr2 | 23 | Surface | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.96
23Sr3 | 23 | Surface | 3
23Dr1 | 23 | Deep | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 8.03
23Dr3 | 23 | Deep | 3
23PRr1 | 23 | Rhizosphere | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.8
23PRr2 | 23 | Rhizosphere | 2
23PRr3 | 23 | Rhizosphere | 3
24Sr1 | 24 | Surface | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.79
24Sr2 | 24 | Surface | 2
24Dr1 | 24 | Deep | 1
24Dr2 | 24 | Deep | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 8.02
24PRr2 | 24 | Rhizosphere | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.97
25Sr1 | 25 | Surface | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.99
25Sr2 | 25 | Surface | 2
25Dr1 | 25 | Deep | 1
25Dr3 | 25 | Deep | 3 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.89
25PRr2 | 25 | Rhizosphere | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.73
26Sr1 | 26 | Surface | 1 | [5 empty] | Depleted
26Sr2 | 26 | Surface | 2 | [5 empty] | Depleted
26Sr3 | 26 | Surface | 3 | [5 empty] | Depleted
26Dr1 | 26 | Deep | 1 | [5 empty] | Depleted
26Dr2 | 26 | Deep | 2 | [5 empty] | Depleted
26Dr3 | 26 | Deep | 3 | [5 empty] | Depleted
26PRr2 | 26 | Rhizosphere | 2 | [5 empty] | Depleted
26PRr3 | 26 | Rhizosphere | 3 | [5 empty] | Depleted
27Sr2 | 27 | Surface | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.92
27Sr3 | 27 | Surface | 3
27Dr1 | 27 | Deep | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.89
27Dr2 | 27 | Deep | 2
27Dr3 | 27 | Deep | 3
27PRr1 | 27 | Rhizosphere | 1
27PRr2 | 27 | Rhizosphere | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.29
28Sr1 | 28 | Surface | 1
28Sr2 | 28 | Surface | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 8
28Sr3 | 28 | Surface | 3
28Dr1 | 28 | Deep | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.98
28Dr3 | 28 | Deep | 3
28PRr1 | 28 | Rhizosphere | 1
28PRr3 | 28 | Rhizosphere | 3 | 2026-07-23 | Yes | 98.1 | 10.04 | 8.01
29Sr1 | 29 | Surface | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.78
29Sr2 | 29 | Surface | 2
29Dr1 | 29 | Deep | 1
29Dr2 | 29 | Deep | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 8.03
29Dr3 | 29 | Deep | 3
29PRr1 | 29 | Rhizosphere | 1
29PRr2 | 29 | Rhizosphere | 2
29PRr3 | 29 | Rhizosphere | 3 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.94
30Sr1 | 30 | Surface | 1
30Sr2 | 30 | Surface | 2
30Sr3 | 30 | Surface | 3 | 2026-07-23 | Yes | 98.1 | 10.04 | 7.93
30Dr1 | 30 | Deep | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 8.1
30PRr1 | 30 | Rhizosphere | 1 | 2026-07-23 | Yes | 98.1 | 10.04 | 8.04
30PRr3 | 30 | Rhizosphere | 3
31Sr2 | 31 | Surface | 2 | 2026-07-23 | Yes | 98.1 | 10.04 | 8.16
31Sr3 | 31 | Surface | 3
31Dr2 | 31 | Deep | 2 | 2026-07-23 | Yes | 98.1 | 10.06 | 8.27
31PRr2 | 31 | Rhizosphere | 2 | 2026-07-23 | Yes | 98.1 | 10.06 | 8
31PRr3 | 31 | Rhizosphere | 3
32Dr3 | 32 | Deep | 3 | 2026-07-23 | Yes | 98.1 | 10.06 | 8.1
32PRr1 | 32 | Rhizosphere | 1
32PRr2 | 32 | Rhizosphere | 2
32PRr3 | 32 | Rhizosphere | 3 | 2026-07-23 | Yes | 98.1 | 10.06 | 8.04
33Sr1 | 33 | Surface | 1
33Sr2 | 33 | Surface | 2 | 2026-07-23 | Yes | 98.1 | 10.06 | 8.08
33Dr1 | 33 | Deep | 1 | 2026-07-23 | Yes | 98.1 | 10.06 | 7.9
33Dr2 | 33 | Deep | 2
33Dr3 | 33 | Deep | 3
33PRr1 | 33 | Rhizosphere | 1
33PRr2 | 33 | Rhizosphere | 2 | 2026-07-23 | Yes | 98.1 | 10.06 | 7.7
34Sr1 | 34 | Surface | 1
34Sr3 | 34 | Surface | 3 | 2026-07-23 | Yes | 98.1 | 10.06 | 7.87
34Dr2 | 34 | Deep | 2
34Dr3 | 34 | Deep | 3 | 2026-07-23 | Yes | 98.1 | 10.06 | 7.81
34PRr1 | 34 | Rhizosphere | 1 | 2026-07-23 | Yes | 98.1 | 10.06 | 7.8
34PRr3 | 34 | Rhizosphere | 3
35Sr1 | 35 | Surface | 1
35Sr2 | 35 | Surface | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.18
35Dr2 | 35 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.92
35Dr3 | 35 | Deep | 3
35PRr1 | 35 | Rhizosphere | 1 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.91
35PRr3 | 35 | Rhizosphere | 3
36Sr1 | 36 | Surface | 1 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.05
36Sr2 | 36 | Surface | 2
36Sr3 | 36 | Surface | 3
36Dr2 | 36 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 8
36Dr3 | 36 | Deep | 3
37Sr1 | 37 | Surface | 1 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.24
37Sr2 | 37 | Surface | 2
37Dr2 | 37 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.22
37Dr3 | 37 | Deep | 3
37PRr2 | 37 | Rhizosphere | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.19
37PRr3 | 37 | Rhizosphere | 3
38Sr2 | 38 | Surface | 2
38Sr3 | 38 | Surface | 3 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.96
38Dr2 | 38 | Deep | 2
38Dr3 | 38 | Deep | 3 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.94
38PRr2 | 38 | Rhizosphere | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.03
38PRr3 | 38 | Rhizosphere | 3
39Sr3 | 39 | Surface | 3 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.92
39Dr2 | 39 | Deep | 2
39Dr3 | 39 | Deep | 3 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.04
39PRr1 | 39 | Rhizosphere | 1
39PRr2 | 39 | Rhizosphere | 2
39PRr3 | 39 | Rhizosphere | 3 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.05
40Dr1 | 40 | Deep | 1 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.96
40Dr2 | 40 | Deep | 2
40PRr2 | 40 | Rhizosphere | 2
40PRr3 | 40 | Rhizosphere | 3 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.06
41Dr1 | 41 | Deep | 1 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.07
41PRr2 | 41 | Rhizosphere | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.98
42Sr1 | 42 | Surface | 1 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.01
42Sr2 | 42 | Surface | 2
42Sr3 | 42 | Surface | 3
42Dr1 | 42 | Deep | 1 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.99
42PRr1 | 42 | Rhizosphere | 1
42PRr2 | 42 | Rhizosphere | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 8.02
43Sr1 | 43 | Surface | 1
43Sr2 | 43 | Surface | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 8.09
43Dr1 | 43 | Deep | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 8.1
43Dr2 | 43 | Deep | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 8.01
43PRr2 | 43 | Rhizosphere | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 8
44Dr2 | 44 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.03 | 7.99
44PRr2 | 44 | Rhizosphere | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8
44PRr3 | 44 | Rhizosphere | 3
45Sr1 | 45 | Surface | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 8.15
45Dr1 | 45 | Deep | 1 | 2026-07-21 | Yes | 97.4 | 10.06 | 8.23
45Dr3 | 45 | Deep | 3
45PRr2 | 45 | Rhizosphere | 2 | 2026-07-21 | Yes | 97.4 | 10.06 | 8.06
46Sr3 | 46 | Surface | 3 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.16
46Dr1 | 46 | Deep | 1 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.2
46PRr1 | 46 | Rhizosphere | 1 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.18
47Sr1 | 47 | Surface | 1 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.14
47Dr2 | 47 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.1
47Dr3 | 47 | Deep | 3 |  |  |  | 10.06
47PRr1 | 47 | Rhizosphere | 1 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.08
48Sr2 | 48 | Surface | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.1
48Dr2 | 48 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8
48PRr3 | 48 | Rhizosphere | 3 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.12
49Sr2 | 49 | Surface | 2 |  |  |  | 10.06 |  | Depleted
49Sr3 | 49 | Surface | 3 |  |  |  | 10.06 |  | Depleted
49Dr2 | 49 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.29
49PRr2 | 49 | Rhizosphere | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.27
50Sr1 | 50 | Surface | 1 |  |  |  | 10.06
50Sr2 | 50 | Surface | 2 |  |  |  | 10.06
50Sr3 | 50 | Surface | 3 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.21
50Dr1 | 50 | Deep | 1 |  |  |  | 10.06
50Dr2 | 50 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.27
50PRr1 | 50 | Rhizosphere | 1 |  |  |  | 10.06
50PRr2 | 50 | Rhizosphere | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.09
51Sr2 | 51 | Surface | 2 |  |  |  | 10.06
51Dr1 | 51 | Deep | 1 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.31
51Dr3 | 51 | Deep | 3
51PRr3 | 51 | Rhizosphere | 3 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.28
52Dr1 | 52 | Deep | 1 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.27
53Sr3 | 53 | Surface | 3 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.42
53Dr2 | 53 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.06 | 8.39
53Dr3 | 53 | Deep | 3
53PRr1 | 53 | Rhizosphere | 1 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.31
53PRr2 | 53 | Rhizosphere | 2
54Sr1 | 54 | Surface | 1 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.15
54Sr2 | 54 | Surface | 2
54Sr3 | 54 | Surface | 3
54Dr1 | 54 | Deep | 1 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.21
54Dr2 | 54 | Deep | 2
54PRr1 | 54 | Rhizosphere | 1 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.17
54PRr2 | 54 | Rhizosphere | 2
54PRr3 | 54 | Rhizosphere | 3
55Sr1 | 55 | Surface | 1 | [5 empty] | Depleted
55Sr3 | 55 | Surface | 3 | [5 empty] | Depleted
55Dr2 | 55 | Deep | 2 | [5 empty] | Depleted
55PRr1 | 55 | Rhizosphere | 1 | [5 empty] | Depleted
55PRr3 | 55 | Rhizosphere | 3 | [5 empty] | Depleted
56Sr1 | 56 | Surface | 1 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.11
56Sr3 | 56 | Surface | 3
56Dr1 | 56 | Deep | 1 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.18
56Dr3 | 56 | Deep | 3
57Sr2 | 57 | Surface | 2 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.43
57Dr3 | 57 | Deep | 3 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.39
57PRr1 | 57 | Rhizosphere | 1
57PRr3 | 57 | Rhizosphere | 3 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.4
58Sr2 | 58 | Surface | 2 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.26
58Sr3 | 58 | Surface | 3
58PRr1 | 58 | Rhizosphere | 1 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.23
59PRr3 | 59 | Rhizosphere | 3 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.45
60Sr3 | 60 | Surface | 3 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.22
60Dr2 | 60 | Deep | 2 | 2026-07-26 | Yes | 97.6 | 10.05 | 8.2
60Dr3 | 60 | Deep | 3

## Trip 2
Sample ID | Site | Compartment | Rep | Date | Calibration | Slope % | pH 10 check | pH reading | Notes
T1Dr1 | 1 | Deep | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.81
T2PRr1 | 2 | Rhizosphere | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.4
T2PRr3 | 2 | Rhizosphere | 3 | 13/7/2026 | Yes | 98 | 10.06 | 7.49
T2Dr2 | 2 | Deep | 2 | 13/7/2026 | Yes | 98 | 10.06 | 7.8
T3Dr2 | 3 | Deep | 2 | 13/7/2026 | Yes | 98 | 10.06 | - | Depleted
T4Sr1 | 4 | Surface | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.79
T4PRr1 | 4 | Rhizosphere | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.7
T4PRr2 | 4 | Rhizosphere | 2 | 13/7/2026 | Yes | 98 | 10.06 | 7.84
T5Sr1 | 5 | Surface | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.75
T5Sr2 | 5 | Surface | 2 | 13/7/2026 | Yes | 98 | 10.06 | 7.75
T5Sr3 | 5 | Surface | 3 | 13/7/2026 | Yes | 98 | 10.06 | 7.82
T5PRr1 | 5 | Rhizosphere | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.59
T6Sr3 | 6 | Surface | 3 | 13/7/2026 | Yes | 98 | 10.06 | - | Depleted
T6Dr1 | 6 | Deep | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.64
T6Dr2 | 6 | Deep | 2 | 13/7/2026 | Yes | 98 | 10.06 | 7.55
T6PRr2 | 6 | Rhizosphere | 2 | 13/7/2026 | Yes | 98 | 10.06 | 7.55
T6PRr3 | 6 | Rhizosphere | 3 | 13/7/2026 | Yes | 98 | 10.06 | 7.57
T7PRr1 | 7 | Rhizosphere | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.62
T7PRr2 | 7 | Rhizosphere | 2 | 13/7/2026 | Yes | 98 | 10.06 | 7.57
T7PRr3 | 7 | Rhizosphere | 3 | 13/7/2026 | Yes | 98 | 10.06 | 7.5
T7Dr3 | 7 | Deep | 3 | 13/7/2026 | Yes | 98 | 10.06 | - | Depleted
T8Sr3 | 8 | Surface | 3 | 13/7/2026 | Yes | 98 | 10.06 | 7.53
T8Dr1 | 8 | Deep | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.91
T8Dr3 | 8 | Deep | 3 | 13/7/2026 | Yes | 98 | 10.06 | 7.87
T8PRr1 | 8 | Rhizosphere | 1 | 13/7/2026 | Yes | 98 | 10.06 | 7.56
T8PRr3 | 8 | Rhizosphere | 3 | 13/7/2026 | Yes | 98 | 10.06 | 7.49

## Trip 3
Sample ID | Site | Compartment | Rep | Date | Calibration | Slope % | pH 10 check | pH reading | Notes
F1Sr1 | 1 | Surface | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.96
F1Sr2 | 1 | Surface | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.97
F1Sr3 | 1 | Surface | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.97
F1Dr1 | 1 | Deep | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.93
F1Dr2 | 1 | Deep | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.99
F1Dr3 | 1 | Deep | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.92
F1PRr1 | 1 | Rhizosphere | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.84
F1PRr2 | 1 | Rhizosphere | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.73
F1PRr3 | 1 | Rhizosphere | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.71
F5Sr1 | 5 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.27
F5Sr2 | 5 | Surface | 2
F5Sr3 | 5 | Surface | 3
F5Dr2 | 5 | Deep | 2 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.2
F5Dr3 | 5 | Deep | 3
F5PRr2 | 5 | Rhizosphere | 2 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.23
F6Sr1 | 6 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.7
F6Sr2 | 6 | Surface | 2
F6Sr3 | 6 | Surface | 3
F6Dr1 | 6 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.74
F6PRr1 | 6 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.69
F6PRr2 | 6 | Rhizosphere | 2
F6PRr3 | 6 | Rhizosphere | 3
F7Sr1 | 7 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.8
F7Sr2 | 7 | Surface | 2
F7Sr3 | 7 | Surface | 3
F7Dr1 | 7 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.84
F7Dr2 | 7 | Deep | 2
F7Dr3 | 7 | Deep | 3
F7PRr1 | 7 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.81
F7PRr2 | 7 | Rhizosphere | 2
F7PRr3 | 7 | Rhizosphere | 3
F8Sr1 | 8 | Surface | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.39
F8Sr2 | 8 | Surface | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.35
F8Sr3 | 8 | Surface | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.35
F8Dr1 | 8 | Deep | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.3
F8Dr2 | 8 | Deep | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.23
F8Dr3 | 8 | Deep | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.36
F8PRr1 | 8 | Rhizosphere | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.68
F8PRr2 | 8 | Rhizosphere | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.76
F8PRr3 | 8 | Rhizosphere | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.68
F9Sr1 | 9 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.69
F9Sr2 | 9 | Surface | 2
F9Sr3 | 9 | Surface | 3
F9Dr1 | 9 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.73
F9Dr2 | 9 | Deep | 2
F9Dr3 | 9 | Deep | 3
F9PRr1 | 9 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.71
F9PRr2 | 9 | Rhizosphere | 2
F9PRr3 | 9 | Rhizosphere | 3
F10Sr1 | 10 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.76
F10Sr2 | 10 | Surface | 2
F10Sr3 | 10 | Surface | 3
F10Dr1 | 10 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.79
F10Dr2 | 10 | Deep | 2
F10Dr3 | 10 | Deep | 3
F10PRr1 | 10 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.7
F10PRr2 | 10 | Rhizosphere | 2
F11Sr1 | 11 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.61
F11Sr2 | 11 | Surface | 2
F11Sr3 | 11 | Surface | 3
F11Dr1 | 11 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.69
F11Dr2 | 11 | Deep | 2
F11Dr3 | 11 | Deep | 3
F11PRr1 | 11 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.65
F11PRr2 | 11 | Rhizosphere | 2
F11PRr3 | 11 | Rhizosphere | 3
F12Sr2 | 12 | Surface | 2 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.05
F12Sr3 | 12 | Surface | 3 | 13/7/2026 | Yes | 98.8 | 10.06 | 8
F12Dr1 | 12 | Deep | 1 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.11
F12Dr2 | 12 | Deep | 2 | 13/7/2026 | Yes | 98.8 | 10.06 | 7.97
F12PRr1 | 12 | Rhizosphere | 1 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.1
F12PRr2 | 12 | Rhizosphere | 2 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.07
F12PRr3 | 12 | Rhizosphere | 3 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.11
F13Sr1 | 13 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.88
F13Sr2 | 13 | Surface | 2
F13Sr3 | 13 | Surface | 3
F13Dr1 | 13 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.9
F13Dr2 | 13 | Deep | 2
F13Dr3 | 13 | Deep | 3
F13PRr1 | 13 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.86
F13PRr2 | 13 | Rhizosphere | 2
F13PRr3 | 13 | Rhizosphere | 3
F14Sr1 | 14 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.97
F14Sr2 | 14 | Surface | 2
F14Sr3 | 14 | Surface | 3
F14Dr1 | 14 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.01
F14Dr2 | 14 | Deep | 2
F14Dr3 | 14 | Deep | 3
F14PRr1 | 14 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.94
F14PRr2 | 14 | Rhizosphere | 2
F14PRr3 | 14 | Rhizosphere | 3
F15Sr1 | 15 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8
F15Sr2 | 15 | Surface | 2
F15Sr3 | 15 | Surface | 3
F15Dr1 | 15 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.02
F15Dr2 | 15 | Deep | 2
F15Dr3 | 15 | Deep | 3
F15PRr1 | 15 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.79
F15PRr2 | 15 | Rhizosphere | 2
F15PRr3 | 15 | Rhizosphere | 3
F16Sr1 | 16 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.96
F16Sr2 | 16 | Surface | 2
F16Sr3 | 16 | Surface | 3
F16Dr1 | 16 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.8
F16Dr2 | 16 | Deep | 2
F16Dr3 | 16 | Deep | 3
F16PRr1 | 16 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.89
F16PRr2 | 16 | Rhizosphere | 2
F16PRr3 | 16 | Rhizosphere | 3
F17Sr1 | 17 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.6
F17Sr2 | 17 | Surface | 2
F17Sr3 | 17 | Surface | 3
F17Dr1 | 17 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.63
F17Dr2 | 17 | Deep | 2
F17Dr3 | 17 | Deep | 3
F17PRr1 | 17 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.58
F17PRr2 | 17 | Rhizosphere | 2
F17PRr3 | 17 | Rhizosphere | 3
F18Sr1 | 18 | Surface | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.96
F18Sr2 | 18 | Surface | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.84
F18Sr3 | 18 | Surface | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.92
F18Dr1 | 18 | Deep | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.59
F18Dr2 | 18 | Deep | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.49
F18Dr3 | 18 | Deep | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.84
F18PRr1 | 18 | Rhizosphere | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.93
F18PRr2 | 18 | Rhizosphere | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.91
F18PRr3 | 18 | Rhizosphere | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.92
F19Sr1 | 19 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.72
F19Sr2 | 19 | Surface | 2
F19Sr3 | 19 | Surface | 3
F19Dr1 | 19 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.77
F19Dr2 | 19 | Deep | 2
F19Dr3 | 19 | Deep | 3
F19PRr1 | 19 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.74
F19PRr2 | 19 | Rhizosphere | 2
F19PRr3 | 19 | Rhizosphere | 3
F20Sr2 | 20 | Surface | 2 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.72
F20Sr3 | 20 | Surface | 3
F20Dr1 | 20 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.77
F20Dr2 | 20 | Deep | 2
F20Dr3 | 20 | Deep | 3
F20PRr1 | 20 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.74
F20PRr2 | 20 | Rhizosphere | 2
F20PRr3 | 20 | Rhizosphere | 3
F21Sr1 | 21 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.53
F21Sr2 | 21 | Surface | 2
F21Sr3 | 21 | Surface | 3
F21Dr1 | 21 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.69
F21Dr2 | 21 | Deep | 2
F21Dr3 | 21 | Deep | 3
F21PRr1 | 21 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.61
F21PRr2 | 21 | Rhizosphere | 2
F21PRr3 | 21 | Rhizosphere | 3
F22Sr1 | 22 | Surface | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.93
F22Sr2 | 22 | Surface | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.96
F22Sr3 | 22 | Surface | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.92
F22Dr1 | 22 | Deep | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.93
F22Dr2 | 22 | Deep | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.82
F22Dr3 | 22 | Deep | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.92
F22PRr1 | 22 | Rhizosphere | 1 | 2026-12-07 | Yes | 100 | 9.95 | 7.93
F22PRr2 | 22 | Rhizosphere | 2 | 2026-12-07 | Yes | 100 | 9.95 | 7.98
F22PRr3 | 22 | Rhizosphere | 3 | 2026-12-07 | Yes | 100 | 9.95 | 7.99
F23Sr1 | 23 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.6
F23Sr2 | 23 | Surface | 2
F23Sr3 | 23 | Surface | 3
F23Dr1 | 23 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.58
F23Dr2 | 23 | Deep | 2
F23Dr3 | 23 | Deep | 3
F23PRr1 | 23 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.54
F23PRr2 | 23 | Rhizosphere | 2
F23PRr3 | 23 | Rhizosphere | 3
F24Sr1 | 24 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.02
F24Sr2 | 24 | Surface | 2
F24Sr3 | 24 | Surface | 3
F24Dr1 | 24 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.03
F24Dr2 | 24 | Deep | 2
F24PRr1 | 24 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.92
F24PRr2 | 24 | Rhizosphere | 2
F24PRr3 | 24 | Rhizosphere | 3
F25Sr1 | 25 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.79
F25Sr2 | 25 | Surface | 2
F25Sr3 | 25 | Surface | 3
F25Dr1 | 25 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.78
F25Dr2 | 25 | Deep | 2
F25Dr3 | 25 | Deep | 3
F25PRr1 | 25 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.7
F25PRr2 | 25 | Rhizosphere | 2
F25PRr3 | 25 | Rhizosphere | 3
F26Sr1 | 26 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.76
F26Sr3 | 26 | Surface | 3
F26Dr1 | 26 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.82
F26Dr3 | 26 | Deep | 3
F26PRr1 | 26 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.75
F26PRr2 | 26 | Rhizosphere | 2
F26PRr3 | 26 | Rhizosphere | 3
F27Sr1 | 27 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.89
F27Sr2 | 27 | Surface | 2
F27Sr3 | 27 | Surface | 3
F27Dr1 | 27 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.95
F27Dr2 | 27 | Deep | 2
F27PRr1 | 27 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.89
F27PRr2 | 27 | Rhizosphere | 2
F27PRr3 | 27 | Rhizosphere | 3
F28Sr1 | 28 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.11
F28Sr2 | 28 | Surface | 2
F28Sr3 | 28 | Surface | 3
F28Dr1 | 28 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.21
F28Dr2 | 28 | Deep | 2
F28Dr3 | 28 | Deep | 3
F28PRr1 | 28 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.03
F28PRr2 | 28 | Rhizosphere | 2
F28PRr3 | 28 | Rhizosphere | 3
F29Sr1 | 29 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 7.93
F29Sr2 | 29 | Surface | 2
F29Sr3 | 29 | Surface | 3
F29Dr1 | 29 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.13
F29Dr2 | 29 | Deep | 2
F29Dr3 | 29 | Deep | 3
F29PRr1 | 29 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.08
F29PRr2 | 29 | Rhizosphere | 2
F29PRr3 | 29 | Rhizosphere | 3
F30Sr1 | 30 | Surface | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.04
F30Sr2 | 30 | Surface | 2
F30Sr3 | 30 | Surface | 3
F30Dr1 | 30 | Deep | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8.16
F30Dr2 | 30 | Deep | 2
F30Dr3 | 30 | Deep | 3
F30PRr1 | 30 | Rhizosphere | 1 | 2026-07-27 | Yes | 97.9 | 10.06 | 8
F30PRr2 | 30 | Rhizosphere | 2
F30PRr3 | 30 | Rhizosphere | 3
F31Sr2 | 31 | Surface | 2 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.98
F31Sr3 | 31 | Surface | 3
F31Dr1 | 31 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.92
F31Dr2 | 31 | Deep | 2
F31PRr1 | 31 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.9
F31PRr2 | 31 | Rhizosphere | 2
F31PRr3 | 31 | Rhizosphere | 3
F32Sr1 | 32 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.82
F32Sr2 | 32 | Surface | 2
F32Sr3 | 32 | Surface | 3
F32Dr1 | 32 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.93
F32Dr2 | 32 | Deep | 2
F32Dr3 | 32 | Deep | 3
F32PRr1 | 32 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.89
F32PRr2 | 32 | Rhizosphere | 2
F32PRr3 | 32 | Rhizosphere | 3
F33Sr1 | 33 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.93
F33Sr2 | 33 | Surface | 2
F33Sr3 | 33 | Surface | 3
F33Dr1 | 33 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 8.01
F33Dr2 | 33 | Deep | 2
F33Dr3 | 33 | Deep | 3
F33PRr2 | 33 | Rhizosphere | 2 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.87
F33PRr3 | 33 | Rhizosphere | 3
F34Sr1 | 34 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.8
F34Sr2 | 34 | Surface | 2
F34Sr3 | 34 | Surface | 3
F34Dr1 | 34 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.79
F34Dr2 | 34 | Deep | 2
F34Dr3 | 34 | Deep | 3
F34PRr1 | 34 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.71
F34PRr2 | 34 | Rhizosphere | 2
F34PRr3 | 34 | Rhizosphere | 3
F35Sr1 | 35 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 8.02
F35Sr2 | 35 | Surface | 2
F35Sr3 | 35 | Surface | 3
F35Dr1 | 35 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 8.17
F35Dr2 | 35 | Deep | 2
F35Dr3 | 35 | Deep | 3
F35PRr1 | 35 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.97
F35PRr2 | 35 | Rhizosphere | 2
F35PRr3 | 35 | Rhizosphere | 3
F36Sr1 | 36 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.88
F36Sr2 | 36 | Surface | 2
F36Dr1 | 36 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 8.07
F36Dr2 | 36 | Deep | 2
F36Dr3 | 36 | Deep | 3
F36PRr1 | 36 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.86
F36PRr2 | 36 | Rhizosphere | 2
F36PRr3 | 36 | Rhizosphere | 3
F37Sr1 | 37 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.98
F37Sr2 | 37 | Surface | 2
F37Sr3 | 37 | Surface | 3
F37Dr1 | 37 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 8.09
F37Dr2 | 37 | Deep | 2
F37Dr3 | 37 | Deep | 3
F37PRr1 | 37 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 8.01
F37PRr2 | 37 | Rhizosphere | 2
F37PRr3 | 37 | Rhizosphere | 3
F38Sr1 | 38 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.83
F38Sr2 | 38 | Surface | 2
F38Sr3 | 38 | Surface | 3
F38Dr2 | 38 | Deep | 2 | 2026-07-28 | Yes | 97.7 | 10.03 | 8.03
F38Dr3 | 38 | Deep | 3
F38PRr1 | 38 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.88
F38PRr2 | 38 | Rhizosphere | 2
F38PRr3 | 38 | Rhizosphere | 3
F39Sr1 | 39 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.03 | 7.91
F39Sr2 | 39 | Surface | 2
F39Sr3 | 39 | Surface | 3
F39Dr2 | 39 | Deep | 2 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.1
F39Dr3 | 39 | Deep | 3
F39PRr1 | 39 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.06
F39PRr3 | 39 | Rhizosphere | 3
F40Sr1 | 40 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.18
F40Sr2 | 40 | Surface | 2
F40Sr3 | 40 | Surface | 3
F40Dr1 | 40 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.1
F40Dr2 | 40 | Deep | 2
F40Dr3 | 40 | Deep | 3
F40PRr1 | 40 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.14
F40PRr2 | 40 | Rhizosphere | 2
F40PRr3 | 40 | Rhizosphere | 3
F41Dr3 | 41 | Deep | 3 | [5 empty] | Depleted
F41PRr1 | 41 | Rhizosphere | 1 | [5 empty] | Depleted
F41PRr2 | 41 | Rhizosphere | 2 | [5 empty] | Depleted
F42Sr1 | 42 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.11
F42Sr2 | 42 | Surface | 2
F42Sr3 | 42 | Surface | 3
F42Dr1 | 42 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.2
F42Dr2 | 42 | Deep | 2
F42Dr3 | 42 | Deep | 3
F42PRr1 | 42 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.09
F42PRr2 | 42 | Rhizosphere | 2
F42PRr3 | 42 | Rhizosphere | 3
F43Sr1 | 43 | Surface | 1 | 13/7/2026 | Yes | 98.8 | 10.05 | 8.14
F43Sr2 | 43 | Surface | 2 | 13/7/2026 | Yes | 98.8 | 10.05 | 8.06
F43Sr3 | 43 | Surface | 3 | 13/7/2026 | Yes | 98.8 | 10.05 | 8.07
F43Dr1 | 43 | Deep | 1 | 13/7/2026 | Yes | 98.8 | 10.05 | 7.97
F43Dr2 | 43 | Deep | 2 | 13/7/2026 | Yes | 98.8 | 10.05 | 8
F43Dr3 | 43 | Deep | 3 | 13/7/2026 | Yes | 98.8 | 10.05 | - | Depleted
F43PRr1 | 43 | Rhizosphere | 1 | 13/7/2026 | Yes | 98.8 | 10.05 | 7.95
F43PRr2 | 43 | Rhizosphere | 2 | 13/7/2026 | Yes | 98.8 | 10.05 | 7.98
F43PRr3 | 43 | Rhizosphere | 3 | 13/7/2026 | Yes | 98.8 | 10.05 | 7.99
F44Sr1 | 44 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.06
F44Sr2 | 44 | Surface | 2
F44Sr3 | 44 | Surface | 3
F44Dr1 | 44 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.01
F44Dr2 | 44 | Deep | 2
F44Dr3 | 44 | Deep | 3
F44PRr1 | 44 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.07
F44PRr2 | 44 | Rhizosphere | 2
F44PRr3 | 44 | Rhizosphere | 3
F45Sr1 | 45 | Surface | 1 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.12
F45Sr2 | 45 | Surface | 2 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.14
F45Sr3 | 45 | Surface | 3 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.17
F45Dr1 | 45 | Deep | 1 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.26
F45Dr2 | 45 | Deep | 2 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.34
F45Dr3 | 45 | Deep | 3 | 13/7/2026 | Yes | 98.8 | 10.06 | 8.33
F45PRr1 | 45 | Rhizosphere | 1 | 13/7/2026 | Yes | 98.8 | 10.06 | 7.87
F45PRr2 | 45 | Rhizosphere | 2 | 13/7/2026 | Yes | 98.8 | 10.06 | 8
F45PRr3 | 45 | Rhizosphere | 3 | 13/7/2026 | Yes | 98.8 | 10.06 | 7.96
F46Sr2 | 46 | Surface | 2 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.05
F46Dr1 | 46 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.13
F47Sr3 | 47 | Surface | 3 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.12
F47Dr1 | 47 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.04
F47Dr3 | 47 | Deep | 3
F47PRr1 | 47 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.16
F47PRr2 | 47 | Rhizosphere | 2
F47PRr3 | 47 | Rhizosphere | 3
F48Sr1 | 48 | Surface | 1 | [5 empty] | Depleted
F48Sr2 | 48 | Surface | 2 | [5 empty] | Depleted
F48Sr3 | 48 | Surface | 3 | [5 empty] | Depleted
F48Dr1 | 48 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.21
F48Dr2 | 48 | Deep | 2 | [5 empty] | Depleted
F48Dr3 | 48 | Deep | 3 | [5 empty] | Depleted
F48PRr1 | 48 | Rhizosphere | 1 | [5 empty] | Depleted
F48PRr2 | 48 | Rhizosphere | 2 | [5 empty] | Depleted
F48PRr3 | 48 | Rhizosphere | 3 | [5 empty] | Depleted
F49Sr1 | 49 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.09
F49Sr2 | 49 | Surface | 2
F49Sr3 | 49 | Surface | 3
F49Dr1 | 49 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 7.94
F49Dr2 | 49 | Deep | 2
F49Dr3 | 49 | Deep | 3
F49PRr1 | 49 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 7.9
F49PRr2 | 49 | Rhizosphere | 2
F49PRr3 | 49 | Rhizosphere | 3
F50Sr1 | 50 | Surface | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.19
F50Sr3 | 50 | Surface | 3
F50Dr1 | 50 | Deep | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.09
F50Dr2 | 50 | Deep | 2
F50Dr3 | 50 | Deep | 3
F50PRr1 | 50 | Rhizosphere | 1 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.2
F50PRr2 | 50 | Rhizosphere | 2
F50PRr3 | 50 | Rhizosphere | 3
F51Sr3 | 51 | Surface | 3 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.43
F51Dr3 | 51 | Deep | 3 | 2026-07-28 | Yes | 97.7 | 10.04 | 8.38
F51PRr1 | 51 | Rhizosphere | 1
F51PRr2 | 51 | Rhizosphere | 2 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.37
F51PRr3 | 51 | Rhizosphere | 3
F52Sr1 | 52 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.41
F52Sr2 | 52 | Surface | 2
F52Sr3 | 52 | Surface | 3
F53Sr3 | 53 | Surface | 3 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.27
F53Dr1 | 53 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.3
F53Dr2 | 53 | Deep | 2
F53Dr3 | 53 | Deep | 3
F53PRr1 | 53 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.2
F54PRr2 | 54 | Rhizosphere | 2 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.1
F55Sr1 | 55 | Surface | 1 | [5 empty] | Depleted
F56Sr2 | 56 | Surface | 2 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.14
F56Sr3 | 56 | Surface | 3
F57Sr3 | 57 | Surface | 3 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.34
F57Dr1 | 57 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.39
F57PRr1 | 57 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.42
F57PRr2 | 57 | Rhizosphere | 2
F57PRr3 | 57 | Rhizosphere | 3
F58Sr1 | 58 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.23
F58Sr2 | 58 | Surface | 2
F58Sr3 | 58 | Surface | 3
F58Dr1 | 58 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.28
F58PRr1 | 58 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.38
F58PRr2 | 58 | Rhizosphere | 2
F58PRr3 | 58 | Rhizosphere | 3
F59Sr1 | 59 | Surface | 1
F59Sr2 | 59 | Surface | 2
F59Sr3 | 59 | Surface | 3 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.49
F59Dr2 | 59 | Deep | 2
F59PRr1 | 59 | Rhizosphere | 1
F60Sr1 | 60 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.53
F60Sr2 | 60 | Surface | 2
F60Sr3 | 60 | Surface | 3
F60PRr1 | 60 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.03 | 8.46
F60PRr2 | 60 | Rhizosphere | 2
F60PRr3 | 60 | Rhizosphere | 3

## Trip 4
Sample ID | Site | Compartment | Rep | Date | Calibration | Slope % | pH 10 check | pH reading | Notes
S1Sr1 | 1 | Surface | 1 | 2026-07-14 | Yes | 95.5 |  | 8.23
S1Dr1 | 1 | Deep | 1 | 2026-07-14 | Yes | 95.5 |  | 8.2
S1PRr1 | 1 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 |  | 8.1
S2Sr1 | 2 | Surface | 1 | 2026-07-14 | Yes | 95.5 |  | 7.91
S2Dr1 | 2 | Deep | 1 | 2026-07-14 | Yes | 95.5 |  | 8.06
S3Sr1 | 3 | Surface | 1 | 2026-07-14 | Yes | 95.5 |  | 7.89
S3Dr1 | 3 | Deep | 1 | 2026-07-14 | Yes | 95.5 |  | 7.97
S3PRr1 | 3 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 |  | 7.96
S4Sr1 | 4 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 8.07
S4Dr1 | 4 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.8
S4PRr1 | 4 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.72
S5Sr1 | 5 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 8.04
S5Dr1 | 5 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.86
S5PRr1 | 5 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.7
S6Sr1 | 6 | Surface | 1 |  |  |  |  | 7.34
S6Dr1 | 6 | Deep | 1 |  |  |  |  | 7.56
S6PRr1 | 6 | Rhizosphere | 1 |  |  |  |  | 7.58
S7Sr1 | 7 | Surface | 1 |  |  |  |  | 7.41
S7Dr1 | 7 | Deep | 1 |  |  |  |  | 7.72
S7PRr1 | 7 | Rhizosphere | 1 |  |  |  |  | 7.33
S8Sr1 | 8 | Surface | 1 |  |  |  |  | 7.32
S8Dr1 | 8 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.64
S8PRr1 | 8 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.57
S9Sr1 | 9 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.65
S9Dr1 | 9 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.56
S9PRr1 | 9 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.61
S10Sr1 | 10 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.53
S10Dr1 | 10 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.4
S10PRr1 | 10 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.51
S11Sr1 | 11 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.83
S11Dr1 | 11 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.9
S11PRr1 | 11 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.79
S12Sr1 | 12 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.07 | 7.84
S12Dr1 | 12 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 8.16
S12PRr1 | 12 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 8.1
S13Sr1 | 13 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 8.04
S13Dr1 | 13 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.86
S13PRr1 | 13 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.93
S14Sr1 | 14 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 8
S14Dr1 | 14 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.9
S14PRr1 | 14 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.96
S15Sr1 | 15 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.88
S15Dr1 | 15 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.77
S15PRr1 | 15 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.85
S16Sr1 | 16 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.7
S16Dr1 | 16 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.64
S16PRr1 | 16 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.87
S17Sr1 | 17 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.9
S17Dr1 | 17 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.71
S17PRr1 | 17 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.99
S18Sr1 | 18 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 8.06
S18Dr1 | 18 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.9
S18PRr1 | 18 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.83
S19Sr1 | 19 | Surface | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.86
S19Dr1 | 19 | Deep | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.7
S19PRr1 | 19 | Rhizosphere | 1 | 2026-07-29 | Yes | 97.6 | 10.05 | 7.64
S20Sr1 | 20 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.97
S20Dr1 | 20 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.75
S20PRr1 | 20 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.71
S21Sr1 | 21 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.81
S21Dr1 | 21 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.93
S21PRr1 | 21 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.84
S22Sr1 | 22 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.76
S22Dr1 | 22 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.66
S22PRr1 | 22 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.74
S23Sr1 | 23 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.81
S23Dr1 | 23 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.95
S23PRr1 | 23 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.91
S24Sr1 | 24 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.76
S24Dr1 | 24 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.63
S24PRr1 | 24 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.71
S25Sr1 | 25 | Surface | 1 |  |  |  |  | 7.4
S25Dr1 | 25 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.69
S25PRr1 | 25 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.78
S26Sr1 | 26 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.61
S26Dr1 | 26 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.7
S26PRr1 | 26 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.75
S27Sr1 | 27 | Surface | 1 |  |  |  |  | 7.55
S27Dr1 | 27 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.77
S27PRr1 | 27 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.05 | 7.82
S28Sr1 | 28 | Surface | 1 | [5 empty] | Missed to prepare
S28Dr2 | 28 | Deep | 2 | 2026-07-30 | Yes | 97 | 10.05 | 7.98
S28PRr1 | 28 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.88
S29Sr1 | 29 | Surface | 1 |  |  |  |  | 7.66
S29Dr1 | 29 | Deep | 1 |  |  |  |  | 7.66
S29PRr1 | 29 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 8
S30Sr1 | 30 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.86
S30Dr1 | 30 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.87
S30PRr1 | 30 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.9
S31Sr1 | 31 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.81
S31Dr1 | 31 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.9
S31PRr1 | 31 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.89
S32Sr1 | 32 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.85
S32Dr1 | 32 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.03 | 8.01
S32PRr1 | 32 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.94
S33Sr1 | 33 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.03 | 8
S33Dr1 | 33 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.93
S33PRr1 | 33 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.88
S34Sr1 | 34 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.87
S34PRr1 | 34 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.86
S35Sr1 | 35 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.98
S35Dr1 | 35 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.03 | 8
S35PRr1 | 35 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.92
S36Sr1 | 36 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.03 | 8.03
S36Dr1 | 36 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.03 | 8
S36PRr1 | 36 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.03 | 7.89
S37Sr1 | 37 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.04
S37Dr1 | 37 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.1
S37PRr1 | 37 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.04 | 7.92
S38Sr1 | 38 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.11
S38Dr1 | 38 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.08
S38PRr1 | 38 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.06
S39Sr1 | 39 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.05
S39Dr1 | 39 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.19
S39PRr1 | 39 | Rhizosphere | 1 | 2026-07-30 | Yes | 97 | 10.04 | 7.92
S40Sr1 | 40 | Surface | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.13
S40Dr1 | 40 | Deep | 1 | 2026-07-30 | Yes | 97 | 10.04 | 8.18
S41Sr1 | 41 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.33
S41Dr1 | 41 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.24
S41PRr1 | 41 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.18
S42Sr1 | 42 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.41
S42Dr1 | 42 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.36
S42PRr1 | 42 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.21
S43Sr1 | 43 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.27
S43Dr1 | 43 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.22
S43PRr1 | 43 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.09
S44Sr1 | 44 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.25
S44Dr1 | 44 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.21
S44PRr1 | 44 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.1
S45Sr1 | 45 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.13
S45Dr1 | 45 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.21
S45PRr1 | 45 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.1
S46Sr1 | 46 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.16
S46Dr1 | 46 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.23
S46PRr1 | 46 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 7.96
S47Sr1 | 47 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.13
S47Dr1 | 47 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.1
S47PRr1 | 47 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.19
S48Sr1 | 48 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.28
S48Dr1 | 48 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.11
S48PRr1 | 48 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.2
S49Sr1 | 49 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.03 | 8.31
S49Dr1 | 49 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.42
S49PRr1 | 49 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.28
S50Sr1 | 50 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.45
S50Dr1 | 50 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.36
S50PRr1 | 50 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.21
S51Sr1 | 51 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 7.81
S51Dr1 | 51 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 7.93
S51PRr1 | 51 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 7.84
S52Sr1 | 52 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.25
S52Dr1 | 52 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.37
S52PRr1 | 52 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.4
S53Sr1 | 53 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.18
S53Dr1 | 53 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.16
S53PRr1 | 53 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.1
S54Sr1 | 54 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.49
S54Dr1 | 54 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.29
S54PRr1 | 54 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.22
S55Sr1 | 55 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.27
S55Dr1 | 55 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.31
S55PRr1 | 55 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.34
S56Sr1 | 56 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.42
S56Dr1 | 56 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.39
S56PRr1 | 56 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.29
S57Sr1 | 57 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.43
S57Dr1 | 57 | Deep | 1 | [5 empty] | Depleted
S57PRr1 | 57 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.06 | 8.32
S58Sr1 | 58 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 8.54
S58Dr1 | 58 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 8.34
S58PRr1 | 58 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 8.24
S59Sr1 | 59 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 8.45
S59Dr1 | 59 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 8.36
S59PRr1 | 59 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 8.47
S60Sr1 | 60 | Surface | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 8.17
S60Dr1 | 60 | Deep | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 7.94
S60PRr1 | 60 | Rhizosphere | 1 | 2026-08-02 | Yes | 97.8 | 10.07 | 9.98

## Trip 5
Sample ID | Site | Compartment | Rep | Date | Calibration | Slope % | pH 10 check | pH reading | Notes
V1Dr1 | 1 | Deep | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.01
V1Dr3 | 1 | Deep | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.05
V1PRr1 | 1 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 8
V1PRr2 | 1 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.92
V1PRr3 | 1 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.86
V2Sr1 | 2 | Surface | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.98
V2Sr2 | 2 | Surface | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.9
V2Dr1 | 2 | Deep | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.11
V2Dr2 | 2 | Deep | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.16
V2PRr1 | 2 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.6
V2PRr2 | 2 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.61
V2PRr3 | 2 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.57
V3PRr1 | 3 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.56
V3PRr2 | 3 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.51
V3PRr3 | 3 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.48
V4Sr1 | 4 | Surface | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.01
V4Sr3 | 4 | Surface | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.96
V4Dr1 | 4 | Deep | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.89
V4Dr2 | 4 | Deep | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.9
V4Dr3 | 4 | Deep | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.94
V4PRr1 | 4 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.94
V4PRr2 | 4 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 8
V4PRr3 | 4 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.91
V5Sr1 | 5 | Surface | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.06
V5Sr3 | 5 | Surface | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.01
V5Dr1 | 5 | Deep | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.01
V5PRr1 | 5 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.03
V5PRr2 | 5 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 8.02
V5PRr3 | 5 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 8
V6Sr1 | 6 | Surface | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.88
V6Sr3 | 6 | Surface | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.85
V6Dr1 | 6 | Deep | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.86
V6PRr1 | 6 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.89
V6PRr2 | 6 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.8
V6PRr3 | 6 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.85
V7Dr1 | 7 | Deep | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.89
V7Dr3 | 7 | Deep | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.8
V7PRr1 | 7 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.85
V7PRr2 | 7 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.8
V7PRr3 | 7 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.04 | 7.81
V8Dr3 | 8 | Deep | 3 | 2026-07-14 | Yes | 95.5 | 10.03 | 7.78
V8PRr1 | 8 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.03 | 7.9
V8PRr2 | 8 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.03 | 7.91
V8PRr3 | 8 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.03 | 7.94
V9PRr1 | 9 | Rhizosphere | 1 | 2026-07-14 | Yes | 95.5 | 10.03 | 7.76
V9PRr2 | 9 | Rhizosphere | 2 | 2026-07-14 | Yes | 95.5 | 10.03 | 7.66
V9PRr3 | 9 | Rhizosphere | 3 | 2026-07-14 | Yes | 95.5 | 10.03 | 7.67
V10Dr1 | 10 | Deep | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.81
V10PRr2 | 10 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.9
V11Sr2 | 11 | Surface | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.85
V11Dr3 | 11 | Deep | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.97
V11PRr2 | 11 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.7
V11PRr3 | 11 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.78
V12Sr1 | 12 | Surface | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.97
V12Sr3 | 12 | Surface | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.9
V12Dr3 | 12 | Deep | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.94
V12PRr1 | 12 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.02
V12PRr2 | 12 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 8
V12PRr3 | 12 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.98
V13Sr3 | 13 | Surface | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.85
V13PRr1 | 13 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.88
V13PRr2 | 13 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.83
V14Sr1 | 14 | Surface | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.09
V14Dr3 | 14 | Deep | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.1
V14PRr1 | 14 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.29
V14PRr2 | 14 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.26
V14PRr3 | 14 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.12
V15Sr1 | 15 | Surface | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.83
V15Sr2 | 15 | Surface | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.79
V15Dr1 | 15 | Deep | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.88
V15Dr2 | 15 | Deep | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.9
V15Dr3 | 15 | Deep | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.86
V15PRr1 | 15 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.09
V15PRr2 | 15 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.14
V15PRr3 | 15 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.99
V16Sr1 | 16 | Surface | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.03 | variant: O
V16Dr1 | 16 | Deep | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.07 | variant: O
V16Dr3 | 16 | Deep | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.06
V16PRr1 | 16 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.26
V16PRr2 | 16 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.28
V16PRr3 | 16 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.2
V17Dr1 | 17 | Deep | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.06
V17Dr2 | 17 | Deep | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.05
V17PRr1 | 17 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.85
V17PRr2 | 17 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.93
V17PRr3 | 17 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.91
V18PRr2 | 18 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.06
V18PRr3 | 18 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.09
V19Sr3 | 19 | Surface | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.8
V19Dr2 | 19 | Deep | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.1
V19Dr3 | 19 | Deep | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 8.12
V19PRr1 | 19 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.97
V19PRr2 | 19 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.95
V19PRr3 | 19 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.93
V20Sr2 | 20 | Surface | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.85
V20PRr1 | 20 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.56
V20PRr2 | 20 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.58
V20PRr3 | 20 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.63
V21Sr3 | 21 | Surface | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.84
V21Dr1 | 21 | Deep | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.7
V21Dr2 | 21 | Deep | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.79
V21PRr1 | 21 | Rhizosphere | 1 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.7
V21PRr2 | 21 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.62
V21PRr3 | 21 | Rhizosphere | 3 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.61
V22PRr2 | 22 | Rhizosphere | 2 | 2026-07-15 | Yes | 97.1 | 10.06 | 7.94
V23Dr1 | 23 | Deep | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 8.17
V23PRr1 | 23 | Rhizosphere | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.7
V23PRr2 | 23 | Rhizosphere | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.67
V24Dr1 | 24 | Deep | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 8.01
V24PRr1 | 24 | Rhizosphere | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.76
V24PRr3 | 24 | Rhizosphere | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.82
V25Sr2 | 25 | Surface | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.7
V25Sr3 | 25 | Surface | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.79
V25Dr3 | 25 | Deep | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.87
V25PRr1 | 25 | Rhizosphere | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.72
V25PRr2 | 25 | Rhizosphere | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.63
V25PRr3 | 25 | Rhizosphere | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.62
V26Sr2 | 26 | Surface | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.67
V26Dr2 | 26 | Deep | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.6
V26Dr3 | 26 | Deep | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.69
V26PRr1 | 26 | Rhizosphere | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.72
V26PRr2 | 26 | Rhizosphere | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.68
V27Dr2 | 27 | Deep | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.74
V27Dr3 | 27 | Deep | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.7
V27PRr1 | 27 | Rhizosphere | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.7
V27PRr2 | 27 | Rhizosphere | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.63
V27PRr3 | 27 | Rhizosphere | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.67
V28Dr1 | 28 | Deep | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.69
V28Dr3 | 28 | Deep | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.7
V28PRr1 | 28 | Rhizosphere | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.66
V28PRr2 | 28 | Rhizosphere | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.6
V28PRr3 | 28 | Rhizosphere | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.57
V29PRr2 | 29 | Rhizosphere | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.73
V29PRr3 | 29 | Rhizosphere | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.65
V30Sr2 | 30 | Surface | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.51
V30PRr1 | 30 | Rhizosphere | 1 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.56
V30PRr2 | 30 | Rhizosphere | 2 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.6
V30PRr3 | 30 | Rhizosphere | 3 | 2026-07-16 | Yes | 97.9 | 10.06 | 7.61
V31Dr1 | 31 | Deep | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.99
V31Dr3 | 31 | Deep | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.88
V32PRr1 | 32 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.65
V32PRr2 | 32 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 7.61
V32PRr3 | 32 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.6
V33PRr1 | 33 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 8.06
V33PRr2 | 33 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 7.98
V33PRr3 | 33 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 8
V34Dr2 | 34 | Deep | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.01
V34PRr1 | 34 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 8.07
V34PRr2 | 34 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.09
V34PRr3 | 34 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 8.11
V35Dr1 | 35 | Deep | 1 | 2026-07-19 | Yes | 98 | 10.06 | 8.1
V35PRr1 | 35 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 8.42
V35PRr2 | 35 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.37
V35PRr3 | 35 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 8.41
V37Dr1 | 37 | Deep | 1 | 2026-07-19 | Yes | 98 | 10.06 | 8.22
V37PRr2 | 37 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.03
V37PRr3 | 37 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.99
V38Sr2 | 38 | Surface | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.14
V38PRr3 | 38 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 8.17
V39Sr2 | 39 | Surface | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.06
V39Dr1 | 39 | Deep | 1 | 2026-07-19 | Yes | 98 | 10.06 | 8.28
V39Dr3 | 39 | Deep | 3 | 2026-07-19 | Yes | 98 | 10.06 | 8.23
V39PRr1 | 39 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 8.18
V39PRr2 | 39 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.2
V39PRr3 | 39 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 8.19
V40Sr3 | 40 | Surface | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.91
V40Dr1 | 40 | Deep | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.99
V40PRr1 | 40 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.89
V40PRr2 | 40 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 7.9
V40PRr3 | 40 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.97
V41Dr3 | 41 | Deep | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.92
V42Sr2 | 42 | Surface | 2 | 2026-07-19 | Yes | 98 | 10.06 | 7.86
V42PRr2 | 42 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.13
V43Dr1 | 43 | Deep | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.9
V43Dr2 | 43 | Deep | 2 | 2026-07-19 | Yes | 98 | 10.06 | 7.85
V43Dr3 | 43 | Deep | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.87
V43PRr1 | 43 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.89
V43PRr2 | 43 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 7.9
V43PRr3 | 43 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.93
V44Dr1 | 44 | Deep | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.83
V44PRr1 | 44 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.95
V44PRr2 | 44 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 7.9
V45Sr3 | 45 | Surface | 3 | 2026-07-19 | Yes | 98 | 10.06 | 8
V45Dr1 | 45 | Deep | 1 | 2026-07-19 | Yes | 98 | 10.06 | 8.12
V45Dr2 | 45 | Deep | 2 | 2026-07-19 | Yes | 98 | 10.06 | 8.15
V45Dr3 | 45 | Deep | 3 | 2026-07-19 | Yes | 98 | 10.06 | 8.2
V45PRr1 | 45 | Rhizosphere | 1 | 2026-07-19 | Yes | 98 | 10.06 | 7.8
V45PRr2 | 45 | Rhizosphere | 2 | 2026-07-19 | Yes | 98 | 10.06 | 7.89
V45PRr3 | 45 | Rhizosphere | 3 | 2026-07-19 | Yes | 98 | 10.06 | 7.9
V46PRr1 | 46 | Rhizosphere | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.06
V46PRr2 | 46 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.09
V46PRr3 | 46 | Rhizosphere | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.05
V47PRr1 | 47 | Rhizosphere | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.24
V47PRr2 | 47 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.28
V47PRr3 | 47 | Rhizosphere | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.2
V48Dr1 | 48 | Deep | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.4
V48Dr2 | 48 | Deep | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.42
V48Dr3 | 48 | Deep | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.37
V48PRr1 | 48 | Rhizosphere | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.31
V49Dr2 | 49 | Deep | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.45
V50PRr2 | 50 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.26
V51PRr2 | 51 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.32
V51PRr3 | 51 | Rhizosphere | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.28
V53PRr2 | 53 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.23
V54PRr1 | 54 | Rhizosphere | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.43
V54PRr2 | 54 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.41
V54PRr3 | 54 | Rhizosphere | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.38
V55Sr2 | 55 | Surface | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.45
V55Sr3 | 55 | Surface | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.4
V55Dr1 | 55 | Deep | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.43
V55PRr1 | 55 | Rhizosphere | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.48
V55PRr2 | 55 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.44
V55PRr3 | 55 | Rhizosphere | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.39
V57PRr2 | 57 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.03
V57PRr3 | 57 | Rhizosphere | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.05
V58PRr1 | 58 | Rhizosphere | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.33
V58PRr2 | 58 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.29
V58PRr3 | 58 | Rhizosphere | 3 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.35
V60PRr1 | 60 | Rhizosphere | 1 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.26
V60PRr2 | 60 | Rhizosphere | 2 | 2026-07-20 | Yes | 97.9 | 10.07 | 8.24

## 7 Jul — P1 micro (Magnus)
7 July 2026 — reference soil mass series, Magnus lab, Probe 1 (microelectrode)
Probe 1: microelectrode, 92% slope, cracked glass. Retained for the record; superseded by 9 July. Fixed 1:2.5 ratio — at constant ratio pH should not vary with mass.
Reference soil — 1 g / 2 g / 5 g, triplicate fresh subsamples
Mass (g) | CaCl₂ (mL) | Rep 1 | Rep 2 | Rep 3 | Mean | Rep SD | Bias vs 5 g
1 | 2.5 | 7.97 | 8 | 7.92 | 7.9633 | 0.040415 | -0.27667
2 | 5 | 8.25 | 8.25 | 8.11 | 8.2033 | 0.080829 | -0.036667
5 | 12.5 | 8.25 | 8.25 | 8.22 | 8.24 | 0.017321 | 0
Reported at the time as “all masses look fine” on replicate scatter alone. Later corrected on two counts: (i) judged against the 5 g value, 1 g reads ~0.28 low; (ii) this probe was subsequently found to have cracked glass and a 92% slope, so the absolute values are not trusted.

## 8 Jul — P2 semi (Magnus)
8 July 2026 — Magnus lab, Probe 2 (semi-micro, 97% slope)
Same soils, same 1:2.5 ratio, 15 mL tubes. Healthy slope, but readings systematically low against known values.
Reference soil — leftover vs freshly prepared CaCl₂
CaCl₂ batch | Mass (g) | Rep 1 | Rep 2 | Rep 3 | Mean | SD | vs 8.03
Leftover (7 Jul) | 1 | 7.28 | 7.51 | 7.51 | 7.4333 | 0.13279 | -0.59667
Leftover (7 Jul) | 1 | 7.69 | 6.89 |  | 7.29 | 0.56569 | -0.74
Fresh (8 Jul) | 1 | 7.56 | 7.58 | 7.61 | 7.5833 | 0.025166 | -0.44667
Fresh (8 Jul) | 5 | 7.74 | 7.72 | 7.78 | 7.7467 | 0.030551 | -0.28333
Fresh CaCl₂ did not restore the expected value — the cleanest reading (fresh, 5 g, SD 0.03) still sat ~0.28 below 8.03. This rules out the electrolyte and points to a probe/lab offset.
Jood's soils re-measured on this probe (vs her earlier 1:2.5 values)
Soil | Mass (g) | Rep 1 | Rep 2 | Rep 3 | Mean | Earlier value | Difference
A8 | 1 | 7.44 | 7.46 | 7.41 | 7.4367
A8 | 2 | 7.38 | 7.32 | 7.33 | 7.3433
A8 | 5 | 7.21 | 7.23 | 7.2 | 7.2133 | 8.21 | -0.99667
A1 | 1 | 7.18 | 7.18 | 7.19 | 7.1833
A1 | 2 | 7.18 | 7.16 | 7.13 | 7.1567
A1 | 5 | 7.09 | 7.05 | 6.99 | 7.0433 | 7.49 | -0.44667

## 9 Jul — P3 ours (valid)
9 July 2026 — reference soil, our lab, Probe 3 (regular, validated)
accumet AB315, pH 7/10 calibration. pH 10 check before / during / after: 10.07, 10.06, 10.08. ATC on, ~20 °C. Fresh 0.01 M CaCl₂.
5 g + 12.5 mL — 50 mL tubes (3 biological replicates, each read twice)
Bio rep | Reading 1 | Reading 2 | Rep mean
1 | 7.97 | 7.94 | 7.955
2 | 8.06 | 8.15 | 8.105
3 | 8.05 | 8 | 8.025
Mean | 8.0267 | 8.03 | 8.0283
SD | 0.049329 | 0.10817 | 0.075056
Established 8.03 ± 0.11 — measured 8.03 ✓ matches
15 mL tubes — same probe (vessel + mass check)
Mass (g) | CaCl₂ (mL) | Rep 1 | Rep 2 | Rep 3 | Mean | SD | Bias vs 5 g
5 | 12.5 | 8.08 |  |  | 8.08 |  | 0
1 | 2.5 | 7.93 | 7.93 | 7.93 | 7.93 | 0 | -0.15
Immersion confirmed for both: inserting the probe displaces liquid and covers the reference junction. 1 g nonetheless reads ~0.15 low with zero replicate scatter — a genuine mass effect, not shallow immersion and not noise.

## Method notes
Method notes and QC rules
Conditions applied across all runs.
Item | Rule
Suspension | 0.01 M CaCl₂, fixed 1:2.5 ratio (mass × 2.5 mL). At constant ratio, pH should not vary with mass.
CaCl₂ prep | 1 M stock; 500 µL stock + 49.5 mL MilliQ = 50 mL of 0.01 M.
Calibration | Two-point, pH 7 and pH 10 (bracketing alkaline soils). pH 4 not used.
Validity criterion | pH 10 read back as an unknown before and after each run; accept if within ~0.1 of nominal.
Slope | Target 95–102%. Recondition below 95%; reject cracked glass.
Temperature | ATC enabled; buffers and samples at the same temperature; recorded each run.
Replicates | Biological = fresh independent scoops (required). Technical re-reads do not substitute — they hide subsampling variation.
Homogenisation | Sieve <2 mm, dry, mix thoroughly; scoop from the well-mixed whole. No fine grinding.
Immersion | Both glass bulb and reference junction must sit in the suspension, clear of the sediment cone.
Reference soil | Run each session; judge against assigned value ± 2 SD (warning) / ± 3 SD (action).

## 1g Reference (log)
In-house reference — establishing the 1 g assigned value
New batch from original bulk. ~10–12 fresh 1 g scoops across ≥3 days and ≥2 calibrations, validated electrode, pH 10 check each session.
Readings (1 g + 2.5 mL, one reading per fresh scoop)
Scoop # | Date | Calibration | Slope % | pH 10 check | Reading (1 g) |  | Notes
1 | 2026-12-07 | Cal-1 | 100 | 9.95 | 7.79 |  | 5 g: 7.95 (12/07)
2 | 2026-12-07 | Cal-1 | 100 | 9.95 | 7.72 |  | 5 g: 8.23 (13/07)
3 | 2026-12-07 | Cal-1 | 100 | 9.95 | 7.75
4 | 2026-12-07 | Cal-1 | 100 | 9.96 | 7.86
5 | 2026-12-07 | Cal-1 | 100 | 9.94 | 7.73
6 | 13/07/2026 | Cal-2 | 98.8 | 10.06 | 7.87
7 | 13/07/2026 | Cal-2 | 98.8 | 10.06 | 7.9
8 | 13/07/2026 | Cal-2 | 98.8 | 10.06 | 8.1
9 | 13/07/2026 | Cal-2 | 98.8 | 10.06 | 7.55
10 | 13/07/2026 | Cal-2 | 98.8 | 10.06 | 7.51
11 | 13/07/2026 | Cal-2 | 98 | 10.06 | 7.59
12 | 13/07/2026 | Cal-2 | 98 | 10.06 | 8.05
13 | 14/07/2026 | Cal-3 | 95.5 | 10.04 | 7.57 |  | 5 g: 8.11 (14/07)
14 | 14/07/2026 | Cal-3 | 95.5 | 10.04 | 7.68
15 | 14/07/2026 | Cal-3 | 95.5 | 10.03 | 7.7
16 | 15/07/2027 | Cal-3 | 97.1 | 10.06 | 7.5
17 | 15/07/2028 | Cal-3 | 97.1 | 10.06 | 7.57
18 | 15/07/2029 | Cal-2 | 97.1 | 10.06 | 8.06
19 | 15/07/2030 | Cal-2 | 97.1 | 10.04 | 7.98
20 | 15/07/2030 | Cal-2 | 97.1 | 10.06 | 8.06
21 | 16/07/2026 | Cal-3 | 97.7 | 10.06 | 8
22 | 1607/2027 | Cal-3 | 97.9 | 10.06 | 7.77
23 | 19/07/2028 | Cal-4 | 98 | 10.06 | 7.69
24 | 1907/2029 | Cal-4 | 98 | 10.06 | 7.95
25 | 20/07/2030 | Cal-5 | 97.9 | 10.07 | 7.98
26 | 20/07/2030 | Cal-5 | 97.9 | 10.07 | 7.67
27 | 21/07/2030 | Cal-6 | 97.4 | 10.06 | 7.58
28 | 21/07/2030 | Cal-6 | 97.4 | 10.06 | 7.78
29 | 22/07/2026 | Cal-7 | 96.1 | 10.06 | 8.06
30 | 22/07/2026 | Cal-7 | 96.1 | 10.07 | 8
31 | 23/07/2026 | Cal-8 | 98.1 | 10.04 | 8.02
32 | 23/07/2026 | Cal-8 | 98.1 | 10.06 | 8.09
33 | 26/07/2026 | Cal-9 | 97.6 | 10.03 | 8.1
34 | 26/07/2026 | Cal-9 | 97.6 | 10.06 | 8.02
35 | 26/07/2026 | Cal-9 | 97.6 | 10.05 | 7.98
36 | 27/07/26 | Cal-10 | 97.9 | 10.06 | 8.11
37 | 27/07/26 | Cal-10 | 97.9 | 10.06 | 7.7
38 | 27/07/2026 | Cal-10 | 97.9 | 10.06 | 7.92
39 | 28/07/2026 | Cal-11 | 97.7 | 10.03 | 7.45
40 | 28/07/2026 | Cal-11 | 97.7 | 10.04 | 8.22
41 | 29/07/2027 | Cal-12 | 97.6 | 10.03 | 8.19
42 | 29/07/2028 | Cal-12 | 97.6 | 10.07 | 8.08
43 | 29/07/2029 | Cal-12 | 97.6 | 10.05 | 8.2
44 | 30/07/2030 | Cal-13 | 97 | 10.05 | 7.67
45 | 30/07/2030 | Cal-13 | 97 | 10.03 | 8.28
46 | 30/07/2030 | Cal-13 | 97 | 10.04 | 8.09
47 | 2026-02-08 | Cal-14 | 97.8 | 10.03 | 8.23
48 | 2026-02-08 | Cal-14 | 97.8 | 10.06 | 7.54
49 | 2026-02-08 | Cal-14 | 97.8 | 10.07 | 7.58
Assigned value and control limits
n (readings) | 28 |  |  | Fill the table as you go; the mean, SD and limits update automatically. Expect the SD to be wider than the old 5 g value (±0.11) because 1 g scoops scatter more — that is honest, not a failure. Aim for ≥10 readings before fixing the value. Optionally run 2–3 scoops at 5 g alongside to record this batch's own 1 g–5 g offset.
Assigned value (mean) | 7.76
SD | 0.18
Warning +2 SD | 8.12
Warning −2 SD | 7.4
Action +3 SD | 8.3
Action −3 SD | 7.22
Day-by-day breakout (watch the within-day scatter)
Session | Scoops | n | Mean | SD | Range | pH 10
Day 1 — 12/07 (Cal-1) | scoops 1–5 | 5 | 7.77 | 0.057 | 0.14 | 9.94–9.96
Day 2 — 13/07 (Cal-2) | scoops 6–12 | 7 | 7.8 | 0.244 | 0.59 | 10
Day 3 — 14/07 (Cal-3) | scoops 13–14 | 2 | 7.63 | 0.078 | 0.11 | 10
Day 4 — 15/07 (Cal-4) | scoops 13–14 | 5 | 7.63 | 0.078 | 0.11 | 10
Day 3 came back tight (range 0.11), like Day 1 — so Day 2's wide scatter looks like an off session, not a heterogeneous batch. Two scoops is thin, though; add 2–3 more Day-3 scoops to be sure before fixing the value. Note between-day drift: the pH 10 check moved 9.95 → 10.06 → 10.04 and the 5 g values 7.95 / 8.23 / 8.11, so some between-day spread is calibration, not soil. ≥3 days and ≥2 calibrations are now met; once a couple more scoops confirm Day-3 tightness, the assigned value can be fixed.

## Replicate test (log)
Replicate-agreement test — can 3 biological replicates → 1?
Three fresh 1 g scoops per compartment, one reading each. If within-compartment SD stays under tolerance across sites, that compartment can move to n=1.
Tolerance (set before measuring): |  |  | 0.15 | ← decide this from what pH difference matters in your analysis. SD at/below this = compartment is a candidate for n=1.
Site ID | Selection rationale (XRF-based) | Compartment | Reps to measure | Rep 1 | Rep 2 | Rep 3 | Mean | SD | Range | ≤ tol?
F8 | lowest Ca (0.53) & lowest Fe — low-carbonate clean sand | Surface | 1,2,3 | 7.39 | 7.35 | 7.35 | 7.3633 | 0.023094 | 0.04 | OK
|  | Deep | 1,2,3 | 7.3 | 7.23 | 7.36 | 7.2967 | 0.065064 | 0.13 | OK
|  | Rhizosphere | 1,2,3 | 7.68 | 7.76 | 7.68 | 7.7067 | 0.046188 | 0.08 | OK
F1 | highest Ca (12.7) — carbonate-rich, high-pH end | Surface | 1,2,3 | 7.96 | 7.97 | 7.97 | 7.9667 | 0.0057735 | 0.01 | OK
|  | Deep | 1,2,3 | 7.93 | 7.99 | 7.92 | 7.9467 | 0.037859 | 0.07 | OK
|  | Rhizosphere | 1,2,3 | 7.84 | 7.73 | 7.71 | 7.76 | 0.07 | 0.13 | OK
F45 | high Ca (7.9) + highest S (0.8) — gypsum-influenced | Surface | 1,2,3 | 8.12 | 8.14 | 8.17 | 7.9433 | 0.066583 | 0.13 | OK
|  | Deep | 1,2,3 | 8.26 | 8.34 | 8.33 | 8.31 | 0.043589 | 0.08 | OK
|  | Rhizosphere | 1,2,3 | 7.87 | 8 | 7.96 | 7.9433 | 0.066583 | 0.13 | OK
F22 | highest Fe (2.25), high Ca — finest / mineral-rich | Surface | 1,2,3 | 7.93 | 7.96 | 7.92 | 7.9367 | 0.020817 | 0.04 | OK
|  | Deep | 1,2,3 | 7.93 | 7.82 | 7.92 | 7.89 | 0.060828 | 0.11 | OK
|  | Rhizosphere | 1,2,3 | 7.93 | 7.98 | 7.99 | 7.9667 | 0.032146 | 0.06 | OK
F18 | near-median Ca (1.4) & Fe — mid-range | Surface | 1,2,3 | 7.96 | 7.84 | 7.92 | 7.9067 | 0.061101 | 0.12 | OK
|  | Deep | 1,2,3 | 7.59 | 7.49 | 7.84 | 7.64 | 0.18028 | 0.35 | review
|  | Rhizosphere | 1,2,3 | 7.93 | 7.91 | 7.92 | 7.92 | 0.01 | 0.02 | OK
F43 | low Fe (0.21), low Ca — second clean-sand check | Surface | 1,2,3 | 8.14 | 8.06 | 8.07 | 8.09 | 0.043589 | 0.08 | OK
|  | Deep | 1,2,3 | 7.97 | 8 | - | 7.985 | 0.021213 | 0.03 | OK
|  | Rhizosphere | 1,2,3 | 7.95 | 7.98 | 7.99 | 7.9733 | 0.020817 | 0.04 | OK
12 | lowest Fe (0.19) — clean quartz sand, likely uniform | Surface | 1,2,3 | 8.14 | 8.05 | 8 | 8.0633 | 0.070946 | 0.14 | OK
|  | Deep | 1,2,3 | 8.11 | 7.97 | 8.04 | 8.04 | 0.07 | 0.14 | OK
|  | Rhizosphere | 1,2,3 | 8.1 | 8.07 | 8.11 | 8.0933 | 0.020817 | 0.04 | OK
Verdict per compartment (fill once several sites are done)
Compartment | Sites tested | Max SD seen | All ≤ tol? | Decision
Surface
Deep
Rhizosphere
Read-off rule: if a compartment's SD stays ≤ tolerance across every site tested → it can go to n=1. If tight at some sites and wide at others → keep n=3 for that compartment. Compartments may differ — selective adoption is fine. A single session is a first look; confirm agreement holds across ≥2 sessions before committing. Run the reference soil each session (see 1g Reference tab).
Site selection basis (Trip 3 / F): sites chosen to satisfy TWO filters — (1) full triplicate sequenced in all three compartments in Trip 3, and (2) spanning carbonate (Ca) and texture (Fe) from the Trip 3 XRF. Ca is the main pH driver; S separates carbonate (low S) from gypsum (high S). F1 = highest Ca; F8/F43 = low-carbonate clean sand; F45 = gypsum-influenced (highest S among full-triplicate sites); F22 = finest/mineral-rich; F18 = mid-range. XRF gives total elements, not pH — it guides site choice only. All six have reps 1,2,3 sequenced, so every compartment gives a real 3-way agreement check.